Type Mark L4-3D (out of production) Manufacturer Mark Products (disappeared) Address Phone E-mail Homepage Category Passive electromagnetic three-component short-period Flat response velocity 1 Hz to >100 Hz (2 Hz version was available) Resolution: NLNM 0.12 to 12 Hz with low-noise preamplifier* within 10 dB of NLNM 0.07 to 35 Hz with low-noise preamplifier* within 20 dB of NLNM 0.05 to 100 Hz with low-noise preamplifier* Operating range large, limited by preamplifier Generator constant 275 Vs/m with 5000 Ohms coil output differential, unlimited Weight 12 kg. Moving mass is 1 kg per component. Size 24 cm dia. , 24 cm high Power - Manual control no mass lock required. Short signal coils and tilt the instrument for transportation. Remote control - Remote diagnostic - Accessories - Typical installation temporary installations, field work Data sheet http://touro.ligo-la.caltech.edu/~ospjeld/HEPI_Documen tation_Website/HEPI_Main_Page/pdfs/L4.pdf Remarks A still widely used short-period seismometer for field work. See publications by Riedesel et al. (BSSA 80,6) and by Rodgers (BSSA 83,2 and 84,1) for information on preamplifier design and noise. The Mark L4-3D is out of production but OYO-Geospace offers a comparable three-component package, the “Seis-Monitor” based on the GS-1 geophone. * based on circuit analysis with op-amp LT1007 as a preamplifier. Low-noise model may be inaccurate above 10 Hz.